## Slide 1: The Initial Experience of New Japanese DCA Device
Etsuo Tsuchikane, MD, PhD
Toyohashi Heart Center
Nagoya Heart Center
Gihu Heart Center

## Slide 2: DCA Catheters History
Designed and developed by John B. Simpson, MD
Clinical application started in 1993 in Japan
History of DCA catheter
  1992 SCA™(Abbott社)
  1995 AtheroCath GTO™(Abbott社)
  1997 AtheroCath BANTAM™(Abbott社)
  2001 FLEXI-CUT®(Abbott社)
  2003 Cypher

## Slide 3: Problems of LMT bifurcation
LADjp: CTO
Pre
GW cross
Stent implantation
Plaque shift and Carina shift
Post
LAD jp : CTO

## Slide 4: Problems of LMT bifurcation
SES3.5x18
Case: 70’s female; LMT
LAD:3.5x18
LCX:3.5x15
POBA4.0x10
T-stent
SES3.0x13
LAD:3.0x13
LCX:3.5x15
Complex stenting

## Slide 5: Problems of LMT bifurcation
Final
8M later
Pre
Complex stenting did not show good outcomes.

## Slide 6
Type D bifurcation
LMT Bifurcation
Notes: I will show you one case. This is a LMT bifurcated lesion involved with diagonal ostium stenosis.

## Slide 7
DCA for LMT~LAD
Notes: First, the directional atherectomy was conducted for distal LMT and ostial LAD. And this is the result after DCA.

## Slide 8
LMT Bifurcation
-LAD Post DCA-

## Slide 9
DCA for LMT~LCx
Notes: Second, DCA was also performed for ostial circumflex artery until an optimal result was obtained as shown here.

## Slide 10
Cypher
Notes: Then we put a drug eluting stent in proximal LAD,

## Slide 11
Cypher
Notes: And in LMT and ostial LAD, as just the crossover stenting.

## Slide 12
KBT
Notes: After that, a high pressure kissing stenting was conducted for both LAD-diagonal and LMT bifurcation.

## Slide 13
Final Angiogram
Notes: This is final angiographic results,

## Slide 14
9Mo Follow-up
Notes: And this is follow-up angiographic results. There was no restenosis.

## Slide 15: PrE Rapamycin Eluting Stent FlExi-CuT Registry (PERFECT Registry)

## Slide 16: Study Design
Multi-center non-randomized prospective registry
DCA debulking prior to Cypher stenting
for ostial or bifurcated lesion
9-month angiographic and 1 yr clinical follow-up
17 Japanese centers

## Slide 17: Endpoints
Primary endpoint
■ Angiographic restenosis (>50% DS) at 9 months
Secondary endpoint
■ MACE, TLR, TVR by 1 yr

## Slide 18: Lesion Location
(n=99)
82%

## Slide 19: DCA Procedural Results
2. Max. cutting pressure: 59.4 ± 34.5 psi
4. Vessels treated by DCA
1) Main branch alone
2) Both main and side branch
3. No of cut: 18.2 ± 12.3 cut
96 (97%)
3 (3%)
1. FlexiCut size L/M: 98/1
(n=99)

## Slide 20: DCA Procedural Results
5. Complications during DCA
1) spasm
2) side branch occlusion
3) no flow
4) perforation
0%
0%
0%
0%
(n=99)

## Slide 21
MB
SB
Debulking with DCA
Stenting to MB
Final
Post DCA
A) Crossover stenting after DCA
With/without KBT or
Ballooning to SB only

## Slide 22
Debulking with DCA
Stenting to MB
Final
MB
SB
Post DCA
B) Non-crossover stenting after DCA

## Slide 23: Stenting Procedural Results
1 Cr. STENT + KBT | 51 % | 50
+ Non KBT | 26 % | 26
+ SB-POBA | 6 % | 6
Non-crossover 1 STENT | 15 % | 15
Crush STENT | 0 % | 0
T STENT | 1 % | 1
Culottes STENT | 1 % | 1
KBT: kissing balloon techique; SB: side branch
(n=99)
82

## Slide 24: In-Hospital Outcomes
1) Death
2) Em-CABG/TLR
3) QMI
4) Non-QMI
5) Complication at access site
0%
0%
0%
2%
0%
(n=99)

## Slide 25: QCU Analysis (Main Branch): Vessel area, Lumen area, Area stenosis
82.3%
55.8%
46.0%
AS

## Slide 26
1yr Follow-up
Results
Fu CAG was performed in 89 pts.
And 96 pts were clinically followed up.
(90% Angiographic Fu rate. 97% Clinical Fu rate.)
No death, no CABG, no MI
TLR for the main branch in 1 patient (1.0%)
TLR for the side branch in 1 patient (1.0%)
(JACC 2007;50:1941-5 )
Notes: And this is follow-up angiographic results. There was no restenosis.
ここの上段3行が表現がいまいちです(加藤が勝手にいじりました) 表題もプロトコールが1年クリニカルなので １Yrにしてます
1.0％になっているのは Clinicalなので分母が96だからです ほかはAngioFUのため分母が89です

## Slide 27
9Mo Follow-up
QCA Results
Binary restenosis rate
Main branch: 1.1% (1/89)
Side branch: 3.4% (3/89)
Fu duration: 264±83 days
(JACC 2007;50:1941-5 )
Notes: And this is follow-up angiographic results. There was no restenosis.

## Slide 28
Follow-up Results
Binary restenosis / TLR rate
(%)
1
1
3
1
4
2
(n=89)
All
(JACC 2007;50:1941-5 )
Notes: And this is follow-up angiographic results. There was no restenosis.

## Slide 29
Follow-up Results
Binary restenosis / TLR rate
(%)
(n=55)
LMT Bifurcation
(JACC 2007;50:1941-5 )
Notes: And this is follow-up angiographic results. There was no restenosis.

## Slide 30
DCA prior to DES stenting was safe and made it possible to avoid complex stenting in bifurcated lesions, which provided a better long-term outcome.
International DCA/DES registry was considered in 2008 for further evaluation.
However, .....
Summary of DCA/DES

## Slide 31: DCA Catheters History
Designed and developed by John B. Simpson, MD
Clinical application started in 1993 in Japan
History of DCA catheter
  1992 SCA™(Abbott社)
  1995 AtheroCath GTO™(Abbott社)
  1997 AtheroCath BANTAM™(Abbott社)
  2001 FLEXI-CUT®(Abbott社)
  2003 Cypher
  2008 Production discontinued
  2015 Revival by NIPRO

## Slide 32: ATHEROCUT® Overview
Guide-wire support
Short cutter w/DLC
Tapered nose-cone with tip
7Fr Guiding Catheter Compatible
6000rpm
Proximal site
Distal site
Motor Drive Unit

## Slide 33: Product Comparison
 | ATHEROCUT® | FLEXI-CUT®
≪Material≫ |  | 
Cutter | DLC; Diamond Like Carbon coating (Vickers hardness：900） | Titan coating (Vickers hardness：550）
≪Specification≫ |  | 
Outer diameter | Φ1.95mm (0.077inch) | Φ2.10mm(0.083inch）
Housing length | 6mm / 9mm | 9mm
Target vessel diameter | 3.0-4.4 mm (3size) | 2.5-4.0 mm (3size)
Nosecorn capacity | 0.032mL | 0.030mL
Guiding Catheter compatibility | ≧ 7Fr ※8Fr recommend | ≧ 8Fr
MDU;rotational speed | 6,000rpm | 3,500rpm

## Slide 34: DLC Cutter -Improved cutting efficiency-
Soft plaque model：cheese
＜ATHEROCUT＞
＜FLEXI-CUT＞
＜ATHEROCUT＞
＜FLEXI-CUT＞
【Functional test of cutting】
Hard model：fluorite (Mohs' scale of hardness4)

## Slide 35: DCA Representative Case From CCT 2016

## Slide 36: CVI Case 60’s male Target Lesion: LMT
Diagnosis: OMI,
Prior intervention :
  ’16.10.02 DCB (Diag.)
Coronary risk factor : Hypertension,Dyslipidemia
eGFR :72.0
Syntax score : 26
Latest CAG findings: ’16.10.07 LMT 75%
LV EF:65% (UCG)

## Slide 37: Pre Angiogram

## Slide 38: IVUS Pre LAD / LCx
LAD
LCx

## Slide 39: DCA 1st session ATHEROCUT® L
1st cut 2atm Lower - far side

## Slide 40: LAD DCA 4th session Max.5atm

## Slide 41: Post DCA

## Slide 42: DCA for LCx 2 session Max. 2atm

## Slide 43: Stent positioning

## Slide 44: Stenting procedure
EES 4.0/18mm 10atm
4.0/15mm 3.0/15mm
KBT&HBT 6atm
SeQuent Please 3.0/20mm 6atm

## Slide 45: Final Angiogram

## Slide 46: Case Summary
Crossover Simple stenting
No plaque/Carina shift

## Slide 47: The Initial Experience of New Japanese DCA Device
Maoto Habara, MD
Etsuo Tsuchikane, MD, PhD
Toyohashi Heart Center
Nagoya Heart Center
Gihu Heart Center
(CCI 2016 in press)

## Slide 48: RESULT
Patients Characteristics
Number | 48
Male | 40 (83%)
Age, years | 66.7±10.8
Prior PCI | 17 (35.4%)
Prior CABG | 1 (2.1%)
HT | 26 (54%)
DM | 15 (31%)
HL | 32 (67%)
H/O Smoking | 16 (33%)
(CCI 2016 in press)

## Slide 49: RESULT
Chart: 
| LMT Bifurcation | LAD Ostial | Cx Ostial | LAD/Diagonal | Cx/OM
 | 16 | 22 | 4 | 6 | 0


Lesion Location
N=48
87%
(CCI 2016 in press)

## Slide 50: RESULT
6 (13%) | 11 (23%) | 1 (2.1%) | 0 (0%) | 2 (4.2%) | 23 (48%) | 5 (10%)
Lesion Location & Bifurcation type (N=48)
LMT Bifurcation | LAD Ostial | LCx Ostial | LAD/ Diagonal | LCx/OM
16 (33%) | 22 (46%) | 4 (8%) | 6 (13%) | 0 (0%)
0
(CCI 2016 in press)

## Slide 51: ACUTE RESULT
Complication during DCA procedure | 
Spasm | 0%
Side branch occlusion | 0%
No flow/ Slow flow | 0%
Perforation | 0%
DCA Procedural Results (N=48)
(CCI 2016 in press)

## Slide 52: ACUTE RESULT
ATHEROCUT® Size L / M | 38 / 11 *
Max. cutting pressure, atm | 5.1±2.5
Number of cuts, cuts | 15.5±8.0
Vessels treated by DCA | 
Main branch alone | 44 (91.7%)
Both main and side branches | 4 (8.3%)
Post DCA %PA, % | 50.1±9.3
※ 1 Left Main bifurcation (1,1,1) case : L size for LM-LAD / M size for LM-LCx
DCA Procedural Results (N=48)
(CCI 2016 in press)

## Slide 53: ACUTE RESULT
Chart: 
| DCA w DES | DCA wo DES
系列 1 | 51.6 | 46.6
列1
列2
Post DCA %PA
51.6±9.2%
46.6±9.3%
P=0.10
(CCI 2016 in press)

## Slide 54: ACUTE RESULT
DCA Procedural Results (N=48)
None stent (DCA alone) | 16 | 33.3%
DCA + DCB | 5 | 10.4%
Stenting | 32 | 66.6%
① Crossover |  | 
1 stent + KBT | 9 | 18.8%
1 stent + non-KBT | 9 | 18.8%
1stent + SB-POBA | 1 | 2.1%
② Non-Crossover | 13 | 27.1%
③ Complex stenting | 0 | 0%
①
②
③
(CCI 2016 in press)

## Slide 55: ACUTE RESULT
Number | 32
Delivery success | 100 (100%)
Stent size, mm | 3.38±0.43
Total length of stent, mm | 20.1±6.7
Max. inflation pressure, atm | 12.7±2.8
Number of Stent / lesion | 1.13±0.34
Stenting Procedural Results (N=32)
(CCI 2016 in press)

## Slide 56: ACUTE RESULT
Device success, % | 100
Procedure success, % | 100
Amount of Contrast medium, ml | 202.4±67.0
Procedure time, min | 112.3±24.7
Fluoroscopy time, min | 40.1±10.7
Procedural Results (N=48)
(CCI 2016 in press)

## Slide 57: ACUTE RESULT
Death | 0%
Em CABG / TLR | 0%
Q MI | 0%
Non-Q MI | 0%
Complication at access site | 0%
In-Hospital Outcome (N=48)
(CCI 2016 in press)

## Slide 58: FOLLOW UP RESULT
Clinical Outcome (N=31: > 6 months)
Fu duration: 375.7 ± 148 days
MACCE | 2 (6.5%)
Cardiac death | 0 (0%)
MI | 0 (0%)
Q MI | 0 (0%)
Non-Q MI | 0 (0%)
TVR | 2 (6.5%)
TLR | 
By PCI | 1 (3.3%)
By CABG | 0 (0%)
Follow up angiography | 19 (61.3%)
Follow up cardiac CT | 1 (3.3%)
Angiographic follow up (N=20: >6 months)
Fu duration: 190.2 ± 86.3 days
Restenosis | 1 (5.0%)
TLR | 1 (5.0%)
(N=48)
(CCI 2016 in press)

## Slide 59: Representative Case DCA + Stent
DCA max10atm 21cuts
%PA 55.6
EES 3.0*15 mm
Post POBA 3.5mm
Pre
Final
F.U
(CCI 2016 in press)

## Slide 60: Representative Case DCA + DCB
DCA for LAD
max4atm 2cuts
DCA for HL
Max10atm 13cuts
Pre
Final
F.U
DCA for LCX
Max1atm 2cuts
DCB 4.0*20 mm
DCB 3.0*20 mm
(CCI 2016 in press)

## Slide 61: Representative Case 40y.o TLR
Pre
Final
F.U
max5atm 17cuts
%PA 41.0%
TLR
MAX5atm 20cuts
%PA 45%
DCB 4.0*20 mm
TLR Final
(CCI 2016 in press)

## Slide 62: SUMMARY
There were no complication during the DCA procedures.
No in-hospital MACE was observed.
DCA could avoid complex stenting in all cases.
In clinical follow up (n=31), TVR rate was 6.5% and TLR rate was 3.3%.
In angiographic follow up (n=20), restenosis rate and TLR rate including both for main and side branches were both 5.0%.
(CCI 2016 in press)

## Slide 63: CONCLUSION
DCA can possibly avoid complex stenting.
Thus, DCA may provide
a good long-term outcome
in patients with major bifurcated lesions.